Study Title: The Effect of Astigmatism on Multifocal Contact Lens Visual Performance – Pilot Study

Participants

Screened n= 15

Enrolled n= 15

Included in Analysis n= 15

Baseline Characteristics

Table 1 - Gender and Age of Study Population

|  | Overall |
| --- | --- |
| N | 15 |
| Gender |  |
| Female | 8 (53.3%) |
| Male | 7 (46.7%) |
| Age |  |
| Mean | 53.1 |
| Std. Dev | 6.9 |

Outcome Measures

Primary Outcome Measure – Visual Performance

Table 2 – Overall Binocular and Monocular Visual Performance (logMAR)

|  | Binocular | Dominant Eye | Non-dominant Eye |
| --- | --- | --- | --- |
| CONTROL |  |  |  |
| N | 15 | 15 | 15 |
| Mean | 0.119 | 0.160 | 0.182 |
| Median | 0.103 | 0.134 | 0.186 |
| TEST |  |  |  |
| N | 15 | 15 | 15 |
| Mean | 0.079 | 0.148 | 0.153 |
| Median | 0.070 | 0.167 | 0.172 |

Secondary Outcome Measure – Visual Satisfaction

Table 3 – Overall Visual Satisfaction (100-point VAS)

|  | Binocular |
| --- | --- |
| CONTROL |  |
| N | 15 |
| Mean | 91.7 |
| Median | 92 |
| TEST |  |
| N | 15 |
| Mean | 89.9 |
| Median | 93.0 |

Adverse Events

There were no adverse events reported in this study.